## Overview
CDP 2.0 Unit Component Status Log
Log containing the current status for each high-level component on each CDP 2.0 unit. Document contains the status, improvement with priority note, general notes on the component status, and current firmware version on the component if applicable.

## Unit A
Component | Parent Module | Status | Priority | Note | Fix By Date | Contact | FW Vers. (in 000')
Seyonic Pipettor | Upper Gantry | 0 |  |  |  | GLC, DB
Seyonic Pipettor Mandrels | Upper Gantry | 0 |  |  |  | GLC, CS
Seyonic Pipettor Tip Eject Bar | Upper Gantry | 0 |  |  |  | GLC, CS
Drip Tray | Upper Gantry | 0 |  |  |  | GLC, AR
Suction Cups | Upper Gantry | 0 |  |  |  | GLC, AR, DB
Suction Cup Air Flow Tubing | Upper Gantry | 0 |  |  |  | GLC
Suction Cup Air Flow Splitters | Upper Gantry | 0 |  |  |  | GLC
Z1 Motor | Upper Gantry | 0 |  |  |  | AR
Z2 Motor | Upper Gantry | 0 |  |  |  | AR
Valve 1 (Tip Eject) | Upper Gantry | 0 |  |  |  | GLC, AR
Valve 2 (Aspirate/Dispense) | Upper Gantry | 0 |  |  |  | GLC, AR
Valve 3 (Suction Cups) | Upper Gantry | 0 |  |  |  | GLC, AR
X-Axis Belt | Upper Gantry | 0
Y-Axis Belt | Upper Gantry | 0
Z1-Axis Belt | Upper Gantry | 0
Z2-Axis Belt | Upper Gantry | 0
X-Axis Mechanical Limit | Upper Gantry | 0 |  |  |  | GLC, AR
Y-Axis Motor | Upper Gantry | 0 |  |  |  | AR
X-Axis Motor | Upper Gantry | 0 |  |  |  | AR
Y-Axis Mechanical Limit | Upper Gantry | 0 |  |  |  | GLC, AR
Z-Axis Mechanical Limit | Upper Gantry | 0 |  |  |  | GLC, AR
Tray AB | Reader | 0 |  |  |  | GLC
Tray CD | Reader | 0 |  |  |  | GLC
Thermocycler A | Reader | 1 | Medium | Thermocycler temperature needs to be calibrated according to Meerstetter | 2022-12-19 | GLC
Thermocycler B | Reader | 1 | Medium | Thermocycler temperature needs to be calibrated according to Meerstetter | 2022-12-19 | GLC
Thermocycler C | Reader | 1 | Medium | Thermocycler temperature needs to be calibrated according to Meerstetter | 2022-12-19 | GLC
Thermocycler D | Reader | 1 | Medium | Thermocycler temperature needs to be calibrated according to Meerstetter | 2022-12-19 | GLC
Thermocycler A Clamp Motor | Reader | 0 |  |  |  | AR
Thermocycler B Clamp Motor | Reader | 0 |  |  |  | AR
Thermocycler C Clamp Motor | Reader | 0 |  |  |  | AR
Thermocycler D Clamp Motor | Reader | 0 |  |  |  | AR
Tray AB Glass | Reader | 0 |  |  |  | GLC, DB
Tray CD Glass | Reader | 0 |  |  |  | GLC, DB
Tray AB Elastomer | Reader | 0 |  |  |  | AJ
Tray CD Elastomer | Reader | 0 |  |  |  | AJ
Hex LED | Reader | 0 |  |  |  | DB
Fam LED | Reader | 0 |  |  |  | DB
Atto590 LED | Reader | 0 |  |  |  | DB
Cy5.5 LED | Reader | 0 |  |  |  | DB
Cy5 LED | Reader | 0 |  |  |  | DB
Alexa405 LED | Reader | 0 |  |  |  | DB
Thermocycler A Fans | Reader | 0 |  |  |  | GLC, AR
Thermocycler B Fans | Reader | 0 |  |  |  | GLC, AR
Thermocycler C Fans | Reader | 0 |  |  |  | GLC, AR
Theromcycler D Fans | Reader | 0 |  |  |  | GLC, AR
Imaging Camera | Reader | 0 |  |  |  | DB
X-Axis Motor | Reader | 0 |  |  |  | AR
Y-Axis Motor | Reader | 0 |  |  |  | AR
Z-Axis Motor | Reader | 0 |  |  |  | AR
Filter Wheel Motor | Reader | 0 |  |  |  | AR
Pre-Amp Thermocycler | Deck Plate | 1 |  | Thermocycler temperature needs to be calibrated according to Meerstetter | 2022-12-19 | GLC
Lid Tray | Deck Plate | 0 |  |  |  | AS, NN
Tip Transfer Tray | Deck Plate | -1 | Medium | Not sufficient for consistent tip pickup requires a redesign |  | AS, NN
Assay Stip | Deck Plate | 0 |  |  |  | AS, NN
Heater/Shaker | Deck Plate | 1 | High | Need to update the Chassis Controller FastAPI code to include the Heater/Shaker | 2022-12-09 | GLC, AR
Mag Separator | Deck Plate | 0 |  |  |  | GLC, AR
Chiller | Deck Plate | 0 | Low | Need to update the Chassis Controller FastAPI code to include the Chiller | 2022-12-19 | GLC, AR
Aux Heater A | Deck Plate | 0 |  |  |  | GLC, AR
Aux Heater B | Deck Plate | 0 |  |  |  | GLC, AR
Aux Heater C | Deck Plate | 0 |  |  |  | GLC, AR
Aux Heater D | Deck Plate | 0 |  |  |  | GLC, AR
Sample Rack A | Deck Plate | 1 | Low | Consumable is not secure resulting in inconsistencies |  | AS, NN
Sample Rack B | Deck Plate | 1 | Low | Consumable is not secure resulting in inconsistencies |  | AS, NN
Sample Rack C | Deck Plate | 1 | Low | Consumable is not secure resulting in inconsistencies |  | AS, NN
Sample Rack D | Deck Plate | 1 | Low | Consumable is not secure resulting in inconsistencies |  | AS, NN
Reagent Cartridge A | Deck Plate | 0 |  |  |  | AS, NN
Reagent Cartridge B | Deck Plate | 0 |  |  |  | AS, NN
Reagent Cartridge C | Deck Plate | 0 |  |  |  | AS, NN
Reagent Cartridge D | Deck Plate | 0 |  |  |  | AS, NN
Tip Tray A | Deck Plate | 0 |  |  |  | AS, NN
Tip Tray B | Deck Plate | 0 |  |  |  | AS, NN
Tip Tray C | Deck Plate | 0 |  |  |  | AS, NN
Tip Tray D | Deck Plate | 0 |  |  |  | AS, NN
Quant Stip A | Deck Plate | 0 |  |  |  | AS, NN
Quant Strip B | Deck Plate | 0 |  |  |  | AS, NN
Quant Strip C | Deck Plate | 0 |  |  |  | AS, NN
Quant Strip D | Deck Plate | 0 |  |  |  | AS, NN
X-Axis Motor PCB | Chassis | 0 |  |  |  | AR | 1.0.0
Y-Axis Motor PCB | Chassis | 0 |  |  |  | AR | 1.0.0
Z-Axis Motor PCB | Chassis | 0 |  |  |  | AR | 1.0.0
Filter Wheel Motor PCB | Chassis | 0 |  |  |  | AR | 1.0.0
Drip Tray Motor PCB | Chassis | 0 |  |  |  | AR | 1.0.0
Z-Axis Motor PCB | Chassis | 0 |  |  |  | AR | 1.0.0
X-Axis Motor PCB | Chassis | 0 |  |  |  | AR | 1.0.0
Valve PCB | Chassis | 0 |  |  |  | AR | 1.0.0
Chassis Power Supply | Chassis | 0 |  |  |  | AR
PC Power Supply | Chassis | 0 |  |  |  | AR
Pre-Amp Thermocycler Power Supply | Chassis | 0 |  |  |  | AR
Thermocycler A Power Supply | Chassis | 0 |  |  |  | AR
Thermocycler B Power Supply | Chassis | 0 |  |  |  | AR
Thermocycler C Power Supply | Chassis | 0 |  |  |  | AR
Thermocycler D Power Supply | Chassis | 0 |  |  |  | AR
Motion Power Supply | Chassis | 0 |  |  |  | AR
Heater/Shaker Chiller Power Supply | Chassis | 0 |  |  |  | AR
Mag Separator Motor PCB | Chassis | 0 |  |  |  | AR | 1.0.0
PC | Chassis | 1 | Low | Currently use an external laptop to control the unit |  | GLC, DB, AR
Chassis Controller PCB | Chassis | 0 |  |  |  | AR | 1.0.0
Relay PCB | Chassis | 0 |  |  |  | AR | 1.0.0
Aux Heater A TEC PCB | Chassis | 0 |  |  |  | GLC, AR | 1.0.0
Aux Heater B TEC PCB | Chassis | 0 |  |  |  | GLC, AR | 1.0.0
Aux Heater C TEC PCB | Chassis | 0 |  |  |  | GLC, AR | 1.0.0
Aux Heater D TEC PCB | Chassis | 0 |  |  |  | GLC, AR | 1.0.0
Pre-Amp Thermocycler TEC PCB | Chassis | 0 |  |  |  | GLC, AR | 1.0.0
Unit LED | Chassis | 0 |  |  |  | AR
Thermocycler A TEC PCB | Chassis | 0 |  |  |  | GLC, AR | 1.0.0
Thermocycler B TEC PCB | Chassis | 0 |  |  |  | GLC, AR | 1.0.0
Thermocycler C TEC PCB | Chassis | 0 |  |  |  | GLC, AR | 1.0.0
Thermocycler D TEC PCB | Chassis | 0 |  |  |  | GLC, AR | 1.0.0
Tray AB Motor PCB | Chassis | 0 |  |  |  | AR | 1.0.0
Tray CD Motor PCB | Chassis | 0 |  |  |  | AR | 1.0.0
Thermocycler Clamp A Motor PCB | Chassis | 0 |  |  |  | AR | 1.0.0
Thermocycler Clamp B Motor PCB | Chassis | 0 |  |  |  | AR | 1.0.0
Thermocycler Clamp C Motor PCB | Chassis | 0 |  |  |  | AR | 1.0.0
Thermocycler Clamp D Motor PCB | Chassis | 0 |  |  |  | AR | 1.0.0
High Pressure Compressor | Chassis | 0 |  |  |  | AR
Intake Fans | Chassis | 0 |  |  |  | AR
Exhaust Fans | Chassis | 0 |  |  |  | AR
Relay 1 | Chassis | 0 |  |  |  | GLC, AR
Relay 2 | Chassis | 0 |  |  |  | GLC, AR
Relay 3 | Chassis | 0 |  |  |  | GLC, AR
Relay 4 | Chassis | 0 |  |  |  | GLC, AR
Relay 5 | Chassis | 0 |  |  |  | GLC, AR
Relay 6 | Chassis | 0 |  |  |  | GLC, AR
Relay 7 | Chassis | 0 |  |  |  | GLC, AR
Relay 8 | Chassis | 0 |  |  |  | GLC, AR
Relay 9 | Chassis | 0 |  |  |  | GLC, AR
Relay 10 | Chassis | 0 |  |  |  | GLC, AR

## Unit B
Component | Parent Module | Status | Priority | Note | Fix By Date | Contact | FW Vers. (in 000')
Seyonic Pipettor | Upper Gantry
Seyonic Pipettor Mandrels | Upper Gantry
Seyonic Pipettor Tip Eject Bar | Upper Gantry
Drip Tray | Upper Gantry
Suction Cups | Upper Gantry
Suction Cup Air Flow Tubing | Upper Gantry
Suction Cup Air Flow Splitters | Upper Gantry
Z1 Motor | Upper Gantry
Z2 Motor | Upper Gantry
Valve 1 (Tip Eject) | Upper Gantry
Valve 2 (Aspirate/Dispense) | Upper Gantry
Valve 3 (Suction Cups) | Upper Gantry
X-Axis Belt | Upper Gantry
Y-Axis Belt | Upper Gantry
Z1-Axis Belt | Upper Gantry
Z2-Axis Belt | Upper Gantry
X-Axis Mechanical Limit | Upper Gantry
Y-Axis Motor | Upper Gantry
X-Axis Motor | Upper Gantry
Y-Axis Mechanical Limit | Upper Gantry
Z-Axis Mechanical Limit | Upper Gantry
Tray AB | Reader
Tray CD | Reader
Thermocycler A | Reader |  | Medium | Thermocycler temperature needs to be calibrated according to Meerstetter | 2022-12-19 | GLC
Thermocycler B | Reader |  | Medium | Thermocycler temperature needs to be calibrated according to Meerstetter | 2022-12-19 | GLC
Thermocycler C | Reader |  | Medium | Thermocycler temperature needs to be calibrated according to Meerstetter | 2022-12-19 | GLC
Thermocycler D | Reader |  | Medium | Thermocycler temperature needs to be calibrated according to Meerstetter | 2022-12-19 | GLC
Thermocycler A Clamp Motor | Reader
Thermocycler B Clamp Motor | Reader
Thermocycler C Clamp Motor | Reader
Thermocycler D Clamp Motor | Reader
Tray AB Glass | Reader
Tray CD Glass | Reader
Tray AB Elastomer | Reader
Tray CD Elastomer | Reader
Hex LED | Reader
Fam LED | Reader
Atto590 LED | Reader
Cy5.5 LED | Reader
Cy5 LED | Reader
Alexa405 LED | Reader
Thermocycler A Fans | Reader
Thermocycler B Fans | Reader
Thermocycler C Fans | Reader
Theromcycler D Fans | Reader
Imaging Camera | Reader
X-Axis Motor | Reader
Y-Axis Motor | Reader
Z-Axis Motor | Reader
Filter Wheel Motor | Reader
Pre-Amp Thermocycler | Deck Plate |  |  | Thermocycler temperature needs to be calibrated according to Meerstetter | 2022-12-19 | GLC
Lid Tray | Deck Plate |  |  |  |  | AS, NN
Tip Transfer Tray | Deck Plate |  | Medium | Not sufficient for consistent tip pickup requires a redesign |  | AS, NN
Assay Stip | Deck Plate |  |  |  |  | AS, NN
Heater/Shaker | Deck Plate |  | High | Need to update the Chassis Controller FastAPI code to include the Heater/Shaker | 2022-12-09 | GLC, AR
Mag Separator | Deck Plate
Chiller | Deck Plate |  | Low | Need to update the Chassis Controller FastAPI code to include the Chiller | 2022-12-19 | GLC, AR
Aux Heater A | Deck Plate
Aux Heater B | Deck Plate
Aux Heater C | Deck Plate
Aux Heater D | Deck Plate
Sample Rack A | Deck Plate |  | Low | Consumable is not secure resulting in inconsistencies |  | AS, NN
Sample Rack B | Deck Plate |  | Low | Consumable is not secure resulting in inconsistencies |  | AS, NN
Sample Rack C | Deck Plate |  | Low | Consumable is not secure resulting in inconsistencies |  | AS, NN
Sample Rack D | Deck Plate |  | Low | Consumable is not secure resulting in inconsistencies |  | AS, NN
Reagent Cartridge A | Deck Plate |  |  |  |  | AS, NN
Reagent Cartridge B | Deck Plate |  |  |  |  | AS, NN
Reagent Cartridge C | Deck Plate |  |  |  |  | AS, NN
Reagent Cartridge D | Deck Plate |  |  |  |  | AS, NN
Tip Tray A | Deck Plate |  |  |  |  | AS, NN
Tip Tray B | Deck Plate |  |  |  |  | AS, NN
Tip Tray C | Deck Plate |  |  |  |  | AS, NN
Tip Tray D | Deck Plate |  |  |  |  | AS, NN
Quant Stip A | Deck Plate |  |  |  |  | AS, NN
Quant Strip B | Deck Plate |  |  |  |  | AS, NN
Quant Strip C | Deck Plate |  |  |  |  | AS, NN
Quant Strip D | Deck Plate |  |  |  |  | AS, NN
X-Axis Motor PCB | Chassis | [5 empty] | 1.0.0
Y-Axis Motor PCB | Chassis | [5 empty] | 1.0.0
Z-Axis Motor PCB | Chassis | [5 empty] | 1.0.0
Filter Wheel Motor PCB | Chassis | [5 empty] | 1.0.0
Drip Tray Motor PCB | Chassis | [5 empty] | 1.0.0
Z-Axis Motor PCB | Chassis | [5 empty] | 1.0.0
X-Axis Motor PCB | Chassis | [5 empty] | 1.0.0
Valve PCB | Chassis | [5 empty] | 1.0.0
Chassis Power Supply | Chassis
PC Power Supply | Chassis
Pre-Amp Thermocycler Power Supply | Chassis
Thermocycler A Power Supply | Chassis
Thermocycler B Power Supply | Chassis
Thermocycler C Power Supply | Chassis
Thermocycler D Power Supply | Chassis
Motion Power Supply | Chassis
Heater/Shaker Chiller Power Supply | Chassis
Mag Separator Motor PCB | Chassis | [5 empty] | 1.0.0
PC | Chassis |  | Low | Currently use an external laptop to control the unit |  | GLC, DB, AR
Chassis Controller PCB | Chassis | [5 empty] | 1.0.0
Relay PCB | Chassis | [5 empty] | 1.0.0
Aux Heater A TEC PCB | Chassis | [5 empty] | 1.0.0
Aux Heater B TEC PCB | Chassis | [5 empty] | 1.0.0
Aux Heater C TEC PCB | Chassis | [5 empty] | 1.0.0
Aux Heater D TEC PCB | Chassis | [5 empty] | 1.0.0
Pre-Amp Thermocycler TEC PCB | Chassis | [5 empty] | 1.0.0
Unit LED | Chassis
Thermocycler A TEC PCB | Chassis | [5 empty] | 1.0.0
Thermocycler B TEC PCB | Chassis | [5 empty] | 1.0.0
Thermocycler C TEC PCB | Chassis | [5 empty] | 1.0.0
Thermocycler D TEC PCB | Chassis | [5 empty] | 1.0.0
Tray AB Motor PCB | Chassis | [5 empty] | 1.0.0
Tray CD Motor PCB | Chassis | [5 empty] | 1.0.0
Thermocycler Clamp A Motor PCB | Chassis | [5 empty] | 1.0.0
Thermocycler Clamp B Motor PCB | Chassis | [5 empty] | 1.0.0
Thermocycler Clamp C Motor PCB | Chassis | [5 empty] | 1.0.0
Thermocycler Clamp D Motor PCB | Chassis | [5 empty] | 1.0.0
High Pressure Compressor | Chassis
Intake Fans | Chassis
Exhaust Fans | Chassis
Relay 1 | Chassis
Relay 2 | Chassis
Relay 3 | Chassis
Relay 4 | Chassis
Relay 5 | Chassis
Relay 6 | Chassis
Relay 7 | Chassis
Relay 8 | Chassis
Relay 9 | Chassis
Relay 10 | Chassis

## Unit C
Component | Parent Module | Status | Priority | Note | Fix By Date | Contact | FW Vers. (in 000')
Seyonic Pipettor | Upper Gantry
Seyonic Pipettor Mandrels | Upper Gantry
Seyonic Pipettor Tip Eject Bar | Upper Gantry
Drip Tray | Upper Gantry
Suction Cups | Upper Gantry
Suction Cup Air Flow Tubing | Upper Gantry
Suction Cup Air Flow Splitters | Upper Gantry
Z1 Motor | Upper Gantry
Z2 Motor | Upper Gantry
Valve 1 (Tip Eject) | Upper Gantry
Valve 2 (Aspirate/Dispense) | Upper Gantry
Valve 3 (Suction Cups) | Upper Gantry
X-Axis Belt | Upper Gantry
Y-Axis Belt | Upper Gantry
Z1-Axis Belt | Upper Gantry
Z2-Axis Belt | Upper Gantry
X-Axis Mechanical Limit | Upper Gantry
Y-Axis Motor | Upper Gantry
X-Axis Motor | Upper Gantry
Y-Axis Mechanical Limit | Upper Gantry
Z-Axis Mechanical Limit | Upper Gantry
Tray AB | Reader
Tray CD | Reader
Thermocycler A | Reader |  | Medium | Thermocycler temperature needs to be calibrated according to Meerstetter | 2022-12-19 | GLC
Thermocycler B | Reader |  | Medium | Thermocycler temperature needs to be calibrated according to Meerstetter | 2022-12-19 | GLC
Thermocycler C | Reader |  | Medium | Thermocycler temperature needs to be calibrated according to Meerstetter | 2022-12-19 | GLC
Thermocycler D | Reader |  | Medium | Thermocycler temperature needs to be calibrated according to Meerstetter | 2022-12-19 | GLC
Thermocycler A Clamp Motor | Reader
Thermocycler B Clamp Motor | Reader
Thermocycler C Clamp Motor | Reader
Thermocycler D Clamp Motor | Reader
Tray AB Glass | Reader
Tray CD Glass | Reader
Tray AB Elastomer | Reader
Tray CD Elastomer | Reader
Hex LED | Reader
Fam LED | Reader
Atto590 LED | Reader
Cy5.5 LED | Reader
Cy5 LED | Reader
Alexa405 LED | Reader
Thermocycler A Fans | Reader
Thermocycler B Fans | Reader
Thermocycler C Fans | Reader
Theromcycler D Fans | Reader
Imaging Camera | Reader
X-Axis Motor | Reader
Y-Axis Motor | Reader
Z-Axis Motor | Reader
Filter Wheel Motor | Reader
Pre-Amp Thermocycler | Deck Plate |  |  | Thermocycler temperature needs to be calibrated according to Meerstetter | 2022-12-19 | GLC
Lid Tray | Deck Plate |  |  |  |  | AS, NN
Tip Transfer Tray | Deck Plate |  | Medium | Not sufficient for consistent tip pickup requires a redesign |  | AS, NN
Assay Stip | Deck Plate |  |  |  |  | AS, NN
Heater/Shaker | Deck Plate |  | High | Need to update the Chassis Controller FastAPI code to include the Heater/Shaker | 2022-12-09 | GLC, AR
Mag Separator | Deck Plate
Chiller | Deck Plate |  | Low | Need to update the Chassis Controller FastAPI code to include the Chiller | 2022-12-19 | GLC, AR
Aux Heater A | Deck Plate
Aux Heater B | Deck Plate
Aux Heater C | Deck Plate
Aux Heater D | Deck Plate
Sample Rack A | Deck Plate |  | Low | Consumable is not secure resulting in inconsistencies |  | AS, NN
Sample Rack B | Deck Plate |  | Low | Consumable is not secure resulting in inconsistencies |  | AS, NN
Sample Rack C | Deck Plate |  | Low | Consumable is not secure resulting in inconsistencies |  | AS, NN
Sample Rack D | Deck Plate |  | Low | Consumable is not secure resulting in inconsistencies |  | AS, NN
Reagent Cartridge A | Deck Plate |  |  |  |  | AS, NN
Reagent Cartridge B | Deck Plate |  |  |  |  | AS, NN
Reagent Cartridge C | Deck Plate |  |  |  |  | AS, NN
Reagent Cartridge D | Deck Plate |  |  |  |  | AS, NN
Tip Tray A | Deck Plate |  |  |  |  | AS, NN
Tip Tray B | Deck Plate |  |  |  |  | AS, NN
Tip Tray C | Deck Plate |  |  |  |  | AS, NN
Tip Tray D | Deck Plate |  |  |  |  | AS, NN
Quant Stip A | Deck Plate |  |  |  |  | AS, NN
Quant Strip B | Deck Plate |  |  |  |  | AS, NN
Quant Strip C | Deck Plate |  |  |  |  | AS, NN
Quant Strip D | Deck Plate |  |  |  |  | AS, NN
X-Axis Motor PCB | Chassis | [5 empty] | 1.0.0
Y-Axis Motor PCB | Chassis | [5 empty] | 1.0.0
Z-Axis Motor PCB | Chassis | [5 empty] | 1.0.0
Filter Wheel Motor PCB | Chassis | [5 empty] | 1.0.0
Drip Tray Motor PCB | Chassis | [5 empty] | 1.0.0
Z-Axis Motor PCB | Chassis | [5 empty] | 1.0.0
X-Axis Motor PCB | Chassis | [5 empty] | 1.0.0
Valve PCB | Chassis | [5 empty] | 1.0.0
Chassis Power Supply | Chassis
PC Power Supply | Chassis
Pre-Amp Thermocycler Power Supply | Chassis
Thermocycler A Power Supply | Chassis
Thermocycler B Power Supply | Chassis
Thermocycler C Power Supply | Chassis
Thermocycler D Power Supply | Chassis
Motion Power Supply | Chassis
Heater/Shaker Chiller Power Supply | Chassis
Mag Separator Motor PCB | Chassis | [5 empty] | 1.0.0
PC | Chassis |  | Low | Currently use an external laptop to control the unit |  | GLC, DB, AR
Chassis Controller PCB | Chassis | [5 empty] | 1.0.0
Relay PCB | Chassis | [5 empty] | 1.0.0
Aux Heater A TEC PCB | Chassis | [5 empty] | 1.0.0
Aux Heater B TEC PCB | Chassis | [5 empty] | 1.0.0
Aux Heater C TEC PCB | Chassis | [5 empty] | 1.0.0
Aux Heater D TEC PCB | Chassis | [5 empty] | 1.0.0
Pre-Amp Thermocycler TEC PCB | Chassis | [5 empty] | 1.0.0
Unit LED | Chassis
Thermocycler A TEC PCB | Chassis | [5 empty] | 1.0.0
Thermocycler B TEC PCB | Chassis | [5 empty] | 1.0.0
Thermocycler C TEC PCB | Chassis | [5 empty] | 1.0.0
Thermocycler D TEC PCB | Chassis | [5 empty] | 1.0.0
Tray AB Motor PCB | Chassis | [5 empty] | 1.0.0
Tray CD Motor PCB | Chassis | [5 empty] | 1.0.0
Thermocycler Clamp A Motor PCB | Chassis | [5 empty] | 1.0.0
Thermocycler Clamp B Motor PCB | Chassis | [5 empty] | 1.0.0
Thermocycler Clamp C Motor PCB | Chassis | [5 empty] | 1.0.0
Thermocycler Clamp D Motor PCB | Chassis | [5 empty] | 1.0.0
High Pressure Compressor | Chassis
Intake Fans | Chassis
Exhaust Fans | Chassis
Relay 1 | Chassis
Relay 2 | Chassis
Relay 3 | Chassis
Relay 4 | Chassis
Relay 5 | Chassis
Relay 6 | Chassis
Relay 7 | Chassis
Relay 8 | Chassis
Relay 9 | Chassis
Relay 10 | Chassis